## NH4-N
Sample ID | Peak Area | [NH4-N] Extract Measured | [NH4-N] Dilution Corrected | [NH4-N] Theoretical | % Recovery | RPD | Units | Est. IDL | Dilution | Anallysis Date | Comments
Cal A | 66.8 | 10.0166 | 10.0166 | 10.0151 | 100.0151 |  | mg/L N | 0.05 | 1 | 2022-08-18 | 202206817A. OM_8-18-2022_01-16-05PM.
Cal B | 27.7 | 3.9956 | 3.9956 | 4.006 | 99.7393 |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal C | 14.4 | 2.0036 | 2.0036 | 2.003 | 100.0272 |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal D | 7.73 | 1.0153 | 1.0153 | 1.0015 | 101.3734 |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal E | 4.29 | 0.50835 | 0.50835 | 0.50075 | 101.5172 |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal F | 2.29 | 0.21451 | 0.21451 | 0.19695 | 108.9144 |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal G | 1.61 | 0.11475 | 0.11475 | 0.10015 | 114.5784 |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal H | 0.519 | -0.045148 | -0.045148 | 0 |  |  | mg/L N | 0.05 | 1 | 2022-08-18
ICV1 | 22.5 | 3.2134 | 3.2134 | 3.29 | 97.6704 |  | mg/L N | 0.05 | 1 | 2022-08-18
ICB1 0.5M K2SO4 | 0.553 | -0.040168 | -0.040168 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
M.Wang P5C | 205 | 33.2682 | 33.2682 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18 | Over-range, diluted and re-ran below
P5D | 205 | 33.2682 | 33.2682 |  |  | 0 | mg/L N | 0.05 | 1 | 2022-08-18 | Over-range, diluted and re-ran below
Blank Not F | 2.74 | 0.28057 | 0.28057 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F Blank P-V? | 3.83 | 0.44071 | 0.44071 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F Blank V? | 4.03 | 0.47011 | 0.47011 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
P4B | 6.93 | 0.8972 | 0.8972 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
P4C | 7.63 | 1.0005 | 1.0005 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
P4D | 9.52 | 1.2798 | 1.2798 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
P5A | 205 | 33.2682 | 33.2682 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18 | Over-range, diluted and re-ran below
P5B | 208 | 33.807 | 33.807 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18 | Over-range, diluted and re-ran below
CCV1 | 29.2 | 4.222 | 4.222 | 4.006 | 105.3921 |  | mg/L N | 0.05 | 1 | 2022-08-18
CCB1 | 0.558 | -0.039436 | -0.039436 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
CCV1 | 29.2 | 3.8519 | 3.8519 | 4.006 | 96.1517 |  | mg/L N | 0.05 | 1 | 2022-08-18
CCB1 | 0.558 | -0.038516 | -0.038516 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
P2D | 7.28 | 0.86718 | 0.86718 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
P3A | 7.21 | 0.85773 | 0.85773 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
P3B | 6.6 | 0.77536 | 0.77536 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
P3C | 7.22 | 0.85908 | 0.85908 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
P4A | 5.84 | 0.67278 | 0.67278 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
P1C | 7.43 | 0.88744 | 0.88744 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
P1D | 7.12 | 0.84557 | 0.84557 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
P2A | 6.11 | 0.70922 | 0.70922 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
P2B | 6.77 | 0.79831 | 0.79831 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
P2C | 5.31 | 0.60128 | 0.60128 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
CCV2 | 29.7 | 3.9205 | 3.9205 | 4.006 | 97.8651 |  | mg/L N | 0.05 | 1 | 2022-08-18
CCB2 | 0.558 | -0.038516 | -0.038516 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
V5B | 11.6 | 1.4516 | 1.4516 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
V5C | 14.2 | 1.8043 | 1.8043 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
V5D | 12.2 | 1.5329 | 1.5329 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
P1A | 6.2 | 0.72136 | 0.72136 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
P1B | 7.75 | 0.93068 | 0.93068 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
V4A | 7.81 | 0.93878 | 0.93878 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
V4B | 12.1 | 1.5194 | 1.5194 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
V4C | 10.6 | 1.3162 | 1.3162 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
V4D | 11.1 | 1.3839 | 1.3839 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
V5A | 10.2 | 1.262 | 1.262 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
CCV3 | 30.2 | 3.9892 | 3.9892 | 4.006 | 99.5792 |  | mg/L N | 0.05 | 1 | 2022-08-18
CCB3 | 0.572 | -0.036635 | -0.036635 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F V2A | 49.4 | 6.6447 | 6.6447 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F V1D | 19.6 | 2.5388 | 2.5388 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F V1C | 19.2 | 2.4843 | 2.4843 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F V1B | 21.6 | 2.8116 | 2.8116 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F V1A | 17.1 | 2.1984 | 2.1984 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F V3B | 33.6 | 4.4567 | 4.4567 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F V3A | 32.7 | 4.3329 | 4.3329 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F V2D | 48.6 | 6.5333 | 6.5333 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F V2C | 44.1 | 5.908 | 5.908 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F V2B | 46.1 | 6.1857 | 6.1857 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
CCV4 | 30.4 | 4.0166 | 4.0166 | 4.006 | 100.265 |  | mg/L N | 0.05 | 1 | 2022-08-18
CCB4 | 0.553 | -0.039188 | -0.039188 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F V4D | 51.1 | 6.8816 | 6.8816 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F V4C | 50.8 | 6.8398 | 6.8398 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F V4B | 54.3 | 7.3283 | 7.3283 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F V4A | 46.1 | 6.1857 | 6.1857 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F V3C | 33.9 | 4.4981 | 4.4981 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F P1A | 7.52 | 0.8996 | 0.8996 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F V5D | 21 | 2.7297 | 2.7297 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F V5C | 20.4 | 2.6479 | 2.6479 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F V5B | 25.4 | 3.331 | 3.331 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F V5A | 17.8 | 2.2936 | 2.2936 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
CCV5 | 30.6 | 4.0441 | 4.0441 | 4.006 | 100.9508 |  | mg/L N | 0.05 | 1 | 2022-08-18
CCB5 | 0.568 | -0.037172 | -0.037172 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F P2B | 12.4 | 1.56 | 1.56 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F P2A | 13.9 | 1.7635 | 1.7635 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F P1D | 8.24 | 0.99689 | 0.99689 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F P1C | 8.41 | 1.0199 | 1.0199 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F P1B | 8.79 | 1.0713 | 1.0713 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F P3C | 13.9 | 1.7635 | 1.7635 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F P3B | 13.5 | 1.7092 | 1.7092 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F P3A | 14.7 | 1.8721 | 1.8721 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F P2D | 13.5 | 1.7092 | 1.7092 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18 | OM_8-18-2022_02-44-06PM.
F P2C | 10.9 | 1.3568 | 1.3568 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
CCV6 | 31.4 | 4.1541 | 4.1541 | 4.006 | 103.6954 |  | mg/L N | 0.05 | 1 | 2022-08-18
CCB6 | 0.623 | -0.02978 | -0.02978 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal A | 73.4 | 10.0158 | 10.0158 | 10.0151 | 100.007 |  | mg/L N | 0.05 | 1 | 2022-08-18 | 20220817B, OM_8-18-2022_02-51-12PM.
Cal B | 30.3 | 4.0029 | 4.0029 | 4.006 | 99.922 |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal C | 15.6 | 1.9944 | 1.9944 | 2.003 | 99.571 |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal D | 8.35 | 1.0118 | 1.0118 | 1.0015 | 101.0243 |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal E | 4.63 | 0.50959 | 0.50959 | 0.50075 | 101.7653 |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal F | 2.44 | 0.2146 | 0.2146 | 0.19695 | 108.9619 |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal G | 1.71 | 0.11638 | 0.11638 | 0.10015 | 116.2053 |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal H | 0.532 | -0.042011 | -0.042011 | 0 |  |  | mg/L N | 0.05 | 1 | 2022-08-18
ICV1 | 24.3 | 3.1805 | 3.1805 | 3.29 | 96.6721 |  | mg/L N | 0.05 | 1 | 2022-08-18
ICB1 0.5M K2SO4 | 0.548 | -0.03986 | -0.03986 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
M.Wang F P5A | 225 | 32.635 | 32.635 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18 | Over-range, diluted and re-ran below
F P4D | 16.6 | 2.1304 | 2.1304 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F P4C | 18.6 | 2.4026 | 2.4026 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F P4B | 17.3 | 2.2256 | 2.2256 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F P4A | 16.3 | 2.0896 | 2.0896 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
V1B | 8.95 | 1.0929 | 1.0929 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
V1A | 6.69 | 0.78751 | 0.78751 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
F P5D | 227 | 32.9487 | 32.9487 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18 | Over-range, diluted and re-ran below
F P5C | 230 | 33.42 | 33.42 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18 | Over-range, diluted and re-ran below
F P5B | 227 | 32.9487 | 32.9487 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18 | Over-range, diluted and re-ran below
CCV1 | 30.6 | 4.0441 | 4.0441 | 4.006 | 100.9508 |  | mg/L N | 0.05 | 1 | 2022-08-18
CCB1 | 0.593 | -0.033812 | -0.033812 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
V2C | 8.2 | 0.99149 | 0.99149 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
C2B | 9.19 | 1.1253 | 1.1253 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
V2A | 8.86 | 1.0807 | 1.0807 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
V1D | 6.49 | 0.76051 | 0.76051 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
V1C | 7.21 | 0.85773 | 0.85773 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
V3C | 5.12 | 0.57566 | 0.57566 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
V3B | 5.07 | 0.56892 | 0.56892 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
V3A | 4.29 | 0.46376 | 0.46376 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
V2D | 9.46 | 1.1619 | 1.1619 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
P5C 1:10 dil. | 41.1 | 5.4922 | 54.9223 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-18
CCV2 | 30.9 | 4.0853 | 4.0853 | 4.006 | 101.9798 |  | mg/L N | 0.05 | 1 | 2022-08-18
CCB2 | 0.591 | -0.034081 | -0.034081 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
P5B 1:10 dil. | 41.3 | 5.5199 | 55.1992 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-18
P5A 1:10 dil. | 38.8 | 5.1741 | 51.7409 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-18
P5D 1:10 dil. | 38.3 | 5.105 | 51.0499 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-18
F P5A 1:10 dil. | 40.7 | 5.4369 | 54.3686 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-18
F P5D 1:10 dil. | 47.7 | 6.4081 | 64.081 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-18
F P5C 1:10 dil. | 44.3 | 5.9357 | 59.3575 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-18
F P5B 1:10 dil. | 36.4 | 4.8427 | 48.4267 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-18
CCV3 | 31.1 | 4.1128 | 4.1128 | 4.006 | 102.666 |  | mg/L N | 0.05 | 1 | 2022-08-18
CCB3 | 0.597 | -0.033275 | -0.033275 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal A | 75.5 | 10.3134 | 10.3134 | 10.0151 | 102.9794 |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal B | 31.1 | 4.1128 | 4.1128 | 4.006 | 102.666 |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal C | 16.1 | 2.0624 | 2.0624 | 2.003 | 102.964 |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal D | 8.54 | 1.0374 | 1.0374 | 1.0015 | 103.589 |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal E | 4.71 | 0.52038 | 0.52038 | 0.50075 | 103.919 |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal F | 2.54 | 0.22806 | 0.22806 | 0.19695 | 115.7957 |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal G | 1.79 | 0.12714 | 0.12714 | 0.10015 | 126.9507 |  | mg/L N | 0.05 | 1 | 2022-08-18
Cal H | 0.613 | -0.031124 | -0.031124 | 0 |  |  | mg/L N | 0.05 | 1 | 2022-08-18

## NO3+NO2-N
Sample ID | Peak Area | [NO3+NO2-N] Extract Measured | [NO3+NO2-N] Dilution Corrected | [NO3+NO2-N] Theoretical | % Recovery | RPD | Units | Est. IDL | Dilution | Anallysis Date | Comments
Cal A | 28.5 | 20.0539 | 20.0539 | 20.0564 | 99.9876 |  | mg/L N | 0.05 | 1 | 2022-08-19 | 20220817A. Data File: OM_8-19-2022_08-05-55AM
Cal B | 13 | 8.0371 | 8.0371 | 8.0225 | 100.1818 |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal C | 3.48 | 1.9684 | 1.9684 | 2.0056 | 98.1413 |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal D | 1.8 | 1.001 | 1.001 | 1.0028 | 99.8214 |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal E | 0.918 | 0.50562 | 0.50562 | 0.50141 | 100.8405 |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal F | 0.38 | 0.20765 | 0.20765 | 0.19721 | 105.2911 |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal G | 0.231 | 0.12568 | 0.12568 | 0.10028 | 125.3309 |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal H | -0.0223 | -0.013089 | -0.013089 | 0 |  |  | mg/L N | 0.05 | 1 | 2022-08-19
ICV1 | 1.6 | 0.88794 | 0.88794 | 0.904 | 98.2234 |  | mg/L N | 0.05 | 1 | 2022-08-19
ICB1 0.5M K2SO4 | 0.0641 | 0.034167 | 0.034167 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
MWang P5C | 60.1 | 52.7497 | 52.7497 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
P5D | 59.1 | 51.5465 | 51.5465 |  |  | 2.3073 | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
Blank Not F | -0.307 | -0.16822 | -0.16822 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Re-ran below
F Blank P? | 0.243 | 0.13228 | 0.13228 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
F Blank V? | 0.253 | 0.13777 | 0.13777 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
P4B | 11.3 | 6.8802 | 6.8802 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
P4C | 12.6 | 7.7621 | 7.7621 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
P4D | 2.55 | 1.429 | 1.429 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
P5A | 59.7 | 52.2671 | 52.2671 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
P5B | 59.2 | 51.6663 | 51.6663 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
CCV1 | 12.6 | 7.7621 | 7.7621 | 8.0225 | 96.753 |  | mg/L N | 0.05 | 1 | 2022-08-19
CCB2 | -0.0226 | -0.013253 | -0.013253 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
Blank Not F Re-run | 0.107 | 0.057661 | 0.057661 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
P2D | 0.626 | 0.3435 | 0.3435 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
P3A | 47 | 37.8607 | 37.8607 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
P3B | 44.5 | 35.2341 | 35.2341 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
P3C | 46.2 | 37.0127 | 37.0127 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
P4A | 9.29 | 5.5534 | 5.5534 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
P1C | 30.7 | 21.974 | 21.974 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
P1D | 34.7 | 25.6016 | 25.6016 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
P2A | 1.02 | 0.56248 | 0.56248 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
P2B | 0.336 | 0.18342 | 0.18342 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
P2C | 0.393 | 0.21481 | 0.21481 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
CCV2 | 12.9 | 7.9682 | 7.9682 | 8.0225 | 99.3226 |  | mg/L N | 0.05 | 1 | 2022-08-19
CCB2 | -0.0206 | -0.01216 | -0.01216 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
V5B | 66 | 60.0729 | 60.0729 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
V5C | 65.4 | 59.3106 | 59.3106 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
V5D | 64.2 | 57.798 | 57.798 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
P1A | 30.4 | 21.709 | 21.709 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
P1B | 31.1 | 22.3288 | 22.3288 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
V4A | 14.4 | 9.0138 | 9.0138 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
V4B | 1.59 | 0.8823 | 0.8823 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
V4C | 2.62 | 1.4693 | 1.4693 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
V4D | 16.9 | 10.8115 | 10.8115 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
V5A | 66.9 | 61.2237 | 61.2237 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
CCV3 | 12.3 | 7.5569 | 7.5569 | 8.0225 | 94.1958 |  | mg/L N | 0.05 | 1 | 2022-08-19
CCB3 | -0.023 | -0.013472 | -0.013472 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
F V2A | 0.552 | 0.30256 | 0.30256 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
F V1D | 34.2 | 25.1385 | 25.1385 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F V1C | 34.1 | 25.0462 | 25.0462 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F V1B | 33 | 24.0383 | 24.0383 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F V1A | 33.5 | 24.4948 | 24.4948 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F V3B | 46.9 | 37.7543 | 37.7543 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F V3A | 48.3 | 39.2538 | 39.2538 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F V2D | 2.4 | 1.3429 | 1.3429 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
F V2C | 0.455 | 0.24899 | 0.24899 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
F V2B | 0.656 | 0.36011 | 0.36011 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
CCV4 | 12.9 | 7.9682 | 7.9682 | 8.0225 | 99.3226 |  | mg/L N | 0.05 | 1 | 2022-08-19
CCB4 | -0.0253 | -0.014729 | -0.014729 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
F V4D | 15.2 | 9.5816 | 9.5816 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
F V4C | 1.86 | 1.035 | 1.035 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
F V4B | 1.07 | 0.59039 | 0.59039 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
F V4A | 13.8 | 8.5926 | 8.5926 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
F V3C | 45.1 | 35.8582 | 35.8582 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F P1A | 31.9 | 23.0437 | 23.0437 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F V5D | 65.8 | 59.8184 | 59.8184 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F V5C | 65.3 | 59.184 | 59.184 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F V5B | 64.4 | 58.049 | 58.049 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F V5A | 66.1 | 60.2003 | 60.2003 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
CCV5 | 12.2 | 7.4887 | 7.4887 | 8.0225 | 93.3461 |  | mg/L N | 0.05 | 1 | 2022-08-19
CCB5 | -0.0338 | -0.019373 | -0.019373 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
F P2B | 0.305 | 0.16636 | 0.16636 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
F P2A | 0.78 | 0.42889 | 0.42889 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
F P1D | 31.5 | 22.6854 | 22.6854 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F P1C | 30.4 | 21.709 | 21.709 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F P1B | 31 | 22.2399 | 22.2399 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F P3C | 46.6 | 37.4358 | 37.4358 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F P3B | 45.1 | 35.8582 | 35.8582 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F P3A | 47.1 | 37.9672 | 37.9672 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F P2D | 0.381 | 0.2082 | 0.2082 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
F P2C | 0.345 | 0.18837 | 0.18837 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
CCV6 | 12.9 | 7.9682 | 7.9682 | 8.0225 | 99.3226 |  | mg/L N | 0.05 | 1 | 2022-08-19
CCB6 | -0.0259 | -0.015057 | -0.015057 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal A | 26.8 | 20.0527 | 20.0527 | 20.0564 | 99.9816 |  | mg/L N | 0.05 | 1 | 2022-08-19 | 20220817B. Data File: OM_8-19-2022_09-38-51AM
Cal B | 12.2 | 8.0436 | 8.0436 | 8.0225 | 100.2621 |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal C | 3.26 | 1.9635 | 1.9635 | 2.0056 | 97.8982 |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal D | 1.68 | 0.9895 | 0.9895 | 1.0028 | 98.6722 |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal E | 0.857 | 0.49414 | 0.49414 | 0.50141 | 98.5496 |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal F | 0.373 | 0.20664 | 0.20664 | 0.19721 | 104.7832 |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal G | 0.218 | 0.11517 | 0.11517 | 0.10028 | 114.8507 |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal H | 0.058 | 0.021059 | 0.021059 | 0 |  |  | mg/L N | 0.05 | 1 | 2022-08-19
ICV1 | 1.51 | 0.88651 | 0.88651 | 0.904 | 98.065 |  | mg/L N | 0.05 | 1 | 2022-08-19
ICB1 0.5M K2SO4 | 0.00161 | -0.012037 | -0.012037 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
MWang F P5A | 56.2 | 52.063 | 52.063 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F P4D | 4.06 | 2.4682 | 2.4682 |  |  | 181.8954 | mg/L N | 0.05 | 1 | 2022-08-19
F P4C | 13.1 | 8.7093 | 8.7093 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
F P4B | 10 | 6.4576 | 6.4576 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
F P4A | 8.69 | 5.5411 | 5.5411 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
V1B | 31.4 | 24.3708 | 24.3708 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
V1A | 30.8 | 23.793 | 23.793 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F P5D | 55.7 | 51.4312 | 51.4312 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F P5C | 56.3 | 52.1898 | 52.1898 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
F P5B | 55.9 | 51.6836 | 51.6836 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
CCV1 | 11.8 | 7.7509 | 7.7509 | 8.0225 | 96.6135 |  | mg/L N | 0.05 | 1 | 2022-08-19
CCB1 | -0.0257 | -0.028052 | -0.028052 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
V2C | 0.683 | 0.39046 | 0.39046 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
V2B | 0.636 | 0.36251 | 0.36251 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
V2A | 0.395 | 0.21965 | 0.21965 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
V1D | 32.2 | 25.1479 | 25.1479 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
V1C | 31.8 | 24.7584 | 24.7584 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
V3C | 43.2 | 36.6187 | 36.6187 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
V3B | 46.3 | 40.1159 | 40.1159 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
V3A | 45.8 | 39.544 | 39.544 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19 | Over calibration range, diluted and re-ran below
V2D | 0.756 | 0.43391 | 0.43391 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
P5C 1:10 | 11.8 | 7.7509 | 77.5087 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
CCV2 | 12.2 | 8.0436 | 8.0436 | 8.0225 | 100.2621 |  | mg/L N | 0.05 | 1 | 2022-08-19
CCB2 | -0.0205 | -0.025003 | -0.025003 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
P5B 1:10 | 11.7 | 7.678 | 76.7799 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
P5A 1:10 | 11.5 | 7.5326 | 75.3261 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
P5D 1:10 | 11.7 | 7.678 | 76.7799 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
F P5A 1:10 | 11.3 | 7.3877 | 73.877 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
F P5D 1:10 | 10.8 | 7.0276 | 70.2756 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
F P5C 1:10 | 11.8 | 7.7509 | 77.5087 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
F P5B 1:10 | 12 | 7.897 | 78.9698 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
P3A 1:10 | 7.79 | 4.9235 | 49.2345 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
P3B 1:10 | 6.49 | 4.0486 | 40.4861 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
P3C 1:10 | 7.38 | 4.6453 | 46.4533 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
CCV3 | 12.3 | 8.1171 | 8.1171 | 8.0225 | 101.178 |  | mg/L N | 0.05 | 1 | 2022-08-19
CCB3 | -0.0213 | -0.025472 | -0.025472 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
P1C 1:10 | 4.24 | 2.5828 | 25.8278 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
P1D 1:10 | 4.47 | 2.7298 | 27.2981 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
P1A 1:10 | 4.15 | 2.5254 | 25.2542 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
P1B 1:10 | 4.15 | 2.5254 | 25.2542 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
V5B 1:10 | 19.4 | 13.6435 | 136.4348 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
V5C 1:10 | 18.4 | 12.8282 | 128.282 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
V5D 1:10 | 20.1 | 14.2214 | 142.2137 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
V5A 1:10 | 19.1 | 13.3976 | 133.9762 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
F P1A 1:10 | 4.31 | 2.6275 | 26.2746 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
F V5D 1:10 | 17.9 | 12.4251 | 124.251 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
CCV4 | 12.2 | 8.0436 | 8.0436 | 8.0225 | 100.2621 |  | mg/L N | 0.05 | 1 | 2022-08-19
CCB4 | -0.0222 | -0.026 | -0.026 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
F V5C 1:10 | 18.5 | 12.9092 | 129.0918 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
F V5B 1:10 | 17.7 | 12.2647 | 122.647 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
F V5A 1:10 | 19.5 | 13.7257 | 137.2567 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
F P1D 1:10 | 4.01 | 2.4364 | 24.3639 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
F P1C 1:10 | 4.5 | 2.749 | 27.4903 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
F P1B 1:10 | 3.88 | 2.3539 | 23.5393 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
F V3B 1:10 | 7.76 | 4.903 | 49.0303 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
F V3A 1:10 | 8 | 5.0667 | 50.6669 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
F P3C 1:10 | 7.71 | 4.869 | 48.6903 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
F P3B 1:10 | 7.15 | 4.4902 | 44.902 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
CCV5 | 12.2 | 8.0436 | 8.0436 | 8.0225 | 100.2621 |  | mg/L N | 0.05 | 1 | 2022-08-19
CCB5 | -0.0233 | -0.026645 | -0.026645 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
F P3A 1:10 | 7.53 | 4.7468 | 47.4685 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
F V3C 1:10 | 7.49 | 4.7197 | 47.1975 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
V1B 1:10 | 4.49 | 2.7426 | 27.4262 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
V1A 1:10 | 4.55 | 2.7811 | 27.811 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
V1D 1:10 | 4.4 | 2.685 | 26.8499 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
V1C 1:10 | 4.53 | 2.7683 | 27.6827 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
V3C 1:10 | 7.01 | 4.3961 | 43.9609 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
V3B 1:10 | 8.68 | 5.5342 | 55.3418 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
V3A 1:10 | 8.18 | 5.1899 | 51.899 |  |  |  | mg/L N | 0.05 | 10 | 2022-08-19
CCV6 | 12.3 | 8.1171 | 8.1171 | 8.0225 | 101.178 |  | mg/L N | 0.05 | 1 | 2022-08-19
CCB6 | -0.0195 | -0.024417 | -0.024417 |  |  |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal A | 26.9 | 20.1438 | 20.1438 | 20.0564 | 100.4361 |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal B | 12.2 | 8.0436 | 8.0436 | 8.0225 | 100.2621 |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal C | 3.26 | 1.9635 | 1.9635 | 2.0056 | 97.8982 |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal D | 1.68 | 0.9895 | 0.9895 | 1.0028 | 98.6722 |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal E | 0.859 | 0.49533 | 0.49533 | 0.50141 | 98.7877 |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal F | 0.376 | 0.20842 | 0.20842 | 0.19721 | 105.6823 |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal G | 0.233 | 0.12401 | 0.12401 | 0.10028 | 123.665 |  | mg/L N | 0.05 | 1 | 2022-08-19
Cal H | 0.00452 | -0.01033 | -0.01033 | 0 |  |  | mg/L N | 0.05 | 1 | 2022-08-19

## Standard Prep
NH4 as N 20220817
Analyte | Std. ID | Stock ID | Stock Conc., mg/L | Stock Conc., mg/g | Flask tare, g | mL Used | g Used | Final Vol., mL | Final Wieght, g | Final Vol. Conc., mg/L | Final Wt. Conc., ug/g | Analyte
NH4 as N | NH4 as N 2022081701 A | NH4-N 2019102101 | 1001.5056 | 1.0025 |  | 1 | 0.9917 | 100 |  | 10.0151 |  | NH4 as N
NH4 as N | NH4 as N 2022081701 B | NH4-N 2019102101 | 1001.5056 | 1.0025 |  | 1 | 0.9962 | 250 |  | 4.006 |  | NH4 as N
NH4 as N | NH4 as N 2022081701 C | NH4-N 2019102101 | 1001.5056 | 1.0025 |  | 0.2 | 0.2004 | 100 |  | 2.003 |  | NH4 as N
NH4 as N | NH4 as N 2022081701 D | NH4-N 2019102101 | 1001.5056 | 1.0025 |  | 0.1 | 0.1004 | 100 |  | 1.0015 |  | NH4 as N
NH4 as N | NH4 as N 2022081701 E | NH4-N 2019102101 | 1001.5056 | 1.0025 |  | 0.05 | 0.0498 | 100 |  | 0.50075 |  | NH4 as N
NH4 as N | NH4 as N 2022081701 F | NH4-N 2019102101 | 1001.5056 | 1.0025 |  | 0.02 | 0.0202 | 101.7 |  | 0.19695 |  | NH4 as N
NH4 as N | NH4 as N 2022081701 G | NH4-N 2019102101 | 1001.5056 | 1.0025 |  | 0.01 | 0.0101 | 100 |  | 0.10015 |  | NH4 as N
NH4 as N | Calibration Blank H | 0.46M K2SO4 | [7 empty] | 0 |  | NH4 as N
Prepared in 0.46M K2SO4 from Fisher Scientific lot # 171510.
NO3 as N 20220817
Analyte | Std. ID | Stock ID | Stock Conc., mg/L | Stock Conc., mg/g | Flask tare, g | mL Used | g Used | Final Vol., mL | Final Wieght, g | Final Vol. Conc., mg/L | Final Wt. Conc., ug/g | Analyte
NO3 as N | NO3 as N 2022081701 A | NO3-N 2019102101 | 1002.8187 | 1.001 |  | 2 | 1.9555 | 100 |  | 20.0564 |  | NO3 as N
NO3 as N | NO3 as N 2022081701 B | NO3-N 2019102101 | 1002.8187 | 1.001 |  | 2 | 1.9819 | 250 |  | 8.0225 |  | NO3 as N
NO3 as N | NO3 as N 2022081701 C | NO3-N 2019102101 | 1002.8187 | 1.001 |  | 0.2 | 0.2002 | 100 |  | 2.0056 |  | NO3 as N
NO3 as N | NO3 as N 2022081701 D | NO3-N 2019102101 | 1002.8187 | 1.001 |  | 0.1 | 0.1001 | 100 |  | 1.0028 |  | NO3 as N
NO3 as N | NO3 as N 2022081701 E | NO3-N 2019102101 | 1002.8187 | 1.001 |  | 0.05 | 0.05 | 100 |  | 0.50141 |  | NO3 as N
NO3 as N | NO3 as N 2022081701 F | NO3-N 2019102101 | 1002.8187 | 1.001 |  | 0.02 | 0.0202 | 101.7 |  | 0.19721 |  | NO3 as N
NO3 as N | NO3 as N 2022081701 G | NO3-N 2019102101 | 1002.8187 | 1.001 |  | 0.01 | 0.0104 | 100 |  | 0.10028 |  | NO3 as N
NO3 as N | Calibration Blank H | 0.46M K2SO4 | [7 empty] | 0 |  | NO3 as N
Prepared in 0.46M K2SO4 from Fisher Scientific lot # 171510.
NO3 as N LCS 20220817
Analyte | Std. ID | Stock ID | Stock Conc., mg/L | Stock Conc., mg/g | Flask tare, g | mL Used | g Used | Final Vol., mL | Final Wieght, g | Final Vol. Conc., mg/L | Final Wt. Conc., ug/g | Analyte
NO3 as N | NO3 as N LCS20220817 | SPEX Lot No. 1-72PWY | 90.36 |  |  | 1 | 0.9977 | 100 |  | 0.9036 |  | NO3 as N
Prepared in 0.46M K2SO4 from Fisher Scientific lot # 171510.
NH4 as N LCS 20220817
Analyte | Std. ID | Stock ID | Stock Conc., mg/L | Stock Conc., mg/g | Flask tare, g | mL Used | g Used | Final Vol., mL | Final Wieght, g | Final Vol. Conc., mg/L | Final Wt. Conc., ug/g | Analyte
NH4 as N | NH4 as N LCS 20220817 | SPEX Lot No.19-168VYY | 328.98 |  |  | 1 | 1.0115 | 100 |  | 3.2898 |  | NH4 as N
Prepared in 0.46M K2SO4 from Fisher Scientific lot # 171510.